## MAIN
| Request Command V1 V2 |  |  |  | Response Command
SRS Command | Opcode | Plen | Payload | CRC16 | Opcode | Plen | Status | Payload | CRC16
SHUTDOWN_SYSTEM | 0xA0 | 0 | --- | 0xXX, 0xXX | 0xA0 | 0 | 0x00 | --- | 0xXX, 0xXX
REBOOT_SYSTEM | 0xA1 | 0 | --- | 0xXX, 0xXX | 0xA1 | 0 | 0x00 | --- | 0xXX, 0xXX
BLE_DISCONNECT_REQUEST | 0xA2 | 0 | --- | 0xXX, 0xXX | 0xA2 | 0 | 0x00 | --- | 0xXX, 0xXX
START_SCHED_THERAPY_SESSION | 0xA3 | 0 1 | VNSb Enable, bool, 1 b | 0xXX, 0xXX | 0xA3 | 0 | 0x00 | --- | 0xXX, 0xXX
END_SCHED_THERAPY_SESSION | 0xA4 | 0 | --- | 0xXX, 0xXX | 0xA4 | 0 | 0x00 | --- | 0xXX, 0xXX
START_MANUAL_THERAPY | 0xA5 | 0 1 | VNSb Enable, bool, 1 b | 0xXX, 0xXX | 0xA5 | 0 | 0x00 | --- | 0xXX, 0xXX
STOP_MANUAL_THERAPY | 0xA6 | 0 | --- | 0xXX, 0xXX | 0xA6 | 0 | 0x00 | --- | 0xXX, 0xXX
GET_IMPEDANCE | 0xA7 | 0 | --- | 0xXX, 0xXX | 0xA7 | 2 | 0x00 | uint16_t ohm(LSB), 2b | 0xXX, 0xXX
MEASURE_BATTERY_VOLTAGE | 0xA8 | 0 | --- | 0xXX, 0xXX | 0xA8 | 2+2 | 0x00 | uint16_t BatA + uint16_t BatB mV(LSB) 2+2b | 0xXX, 0xXX
ERASE_LOG | 0xAF | 0 | --- | 0xXX, 0xXX | 0xAF | 0 | 0x00 | --- | 0xXX, 0xXX
START_MANUAL_4X_THERAPY | 0xB2 | 9 | CH1~CH8+ENCL Status, uint8_t, 9 b 0x00: Cathode 0x01: Anode 0xFF: Disable | 0xXX, 0xXX | 0xB2 | 0 | 0x00 | --- | 0xXX, 0xXX
[7 empty] | 0 | STATUS_SUCCESS
[7 empty] | 1 | STATUS_INVALID
[7 empty] | F0 | STATUS_CRC_ERR
[7 empty] | F1 | STATUS_PAYLOAD_LEN_ERR
[7 empty] | F2 | STATUS_OPCODE_ERR
[7 empty] | F3 | STATUS_USER_CLASS_ERR

## GET_SENSOR
Request Command |  |  |  |  | Response Command
Opcode | Plen | Payload(Sensor ID, 1byte) | Payload | CRC16 | Opcode | Plen | Status | P-Serial | P-mV | CRC16
0xA9 | 1 | 0x00, IDLE | --- | 0xXX, 0xXX | 0xA9 | 0 | 0x00 | --- | --- | 0xXX, 0xXX
0xA9 | 1 | 0x01, ECG_HR, 256Hz | --- | 0xXX, 0xXX | 0xA9 | 1 + 200 | 0x00 | 0~99, Serial N, 1b | uint16_t, LSB, 100 x mV, 200b | 0xXX, 0xXX
0xA9 | 1 | 0x02, ECG_RR, 256Hz | --- | 0xXX, 0xXX | 0xA9 | 1 + 200 | 0x00 | 0~99, Serial N, 1b | uint16_t, LSB, 100 x mV, 200b | 0xXX, 0xXX
0xA9 | 1 | 0x03, EMG1, 5000Hz | --- | 0xXX, 0xXX | 0xA9 | 1 + 200 | 0x00 | 0~99, Serial N, 1b | uint16_t, LSB, 100 x mV, 200b | 0xXX, 0xXX
0xA9 | 1 | 0x04, EMG2, 5000Hz | --- | 0xXX, 0xXX | 0xA9 | 1 + 200 | 0x00 | 0~99, Serial N, 1b | uint16_t, LSB, 100 x mV, 200b | 0xXX, 0xXX
0xA9 | 1 + 4 | 0x01, ECG_HR | float, Freq, LSB, 1.5 ~ 6553.5Hz, 4b | 0xXX, 0xXX | 0xA9 | 1 + 200 | 0x00 | 0~99, Serial N, 1b | uint16_t, LSB, 100 x mV, 200b | 0xXX, 0xXX
0xA9 | 1 + 4 | 0x02, ECG_RR | float, Freq, LSB, 1.5 ~ 6553.5Hz, 4b | 0xXX, 0xXX | 0xA9 | 1 + 200 | 0x00 | 0~99, Serial N, 1b | uint16_t, LSB, 100 x mV, 200b | 0xXX, 0xXX
0xA9 | 1 + 4 | 0x03, EMG1 | float, Freq, LSB, 1.5 ~ 6553.5Hz, 4b | 0xXX, 0xXX | 0xA9 | 1 + 200 | 0x00 | 0~99, Serial N, 1b | uint16_t, LSB, 100 x mV, 200b | 0xXX, 0xXX
0xA9 | 1 + 4 | 0x04, EMG2 | float, Freq, LSB, 1.5 ~ 6553.5Hz, 4b | 0xXX, 0xXX | 0xA9 | 1 + 200 | 0x00 | 0~99, Serial N, 1b | uint16_t, LSB, 100 x mV, 200b | 0xXX, 0xXX

## GET_ACC
Request Command |  |  |  | Response Command
Opcode | Plen | P-DevID, 1b | CRC16 | Opcode | Plen | Status | P-DevID,1b | P-Tstamp,4b | P-Index, 2b | P-XYZdata, 192b | CRC16
0x48 | 1 | ID: 0x18, 0x19 | 0xXX, 0xXX | 0x48 | 1+4+2+upto192 | 0x00 | ID: 0x18 or 0x19 | msec, uint32_t(LSB) | Index, uint16_t(LSB) | XYZ data, uint16_t(LSB) x 3 axis x upto32 samp | 0xXX, 0xXX
|  |  |  | 0 | 1 | 2 | 3 | 4--7 | 8--9

## START_ACC
Request Command |  |  |  |  | Response Command
Opcode | Plen | P-FreqSamp, 1b | P-Mode+DevID, 1b | CRC16 | Opcode | Plen | Status | P-DevID, 1b | P-Xdata, 2b | P-Ydata, 2b | P-Zdata, 2b | CRC16
0x47 | 2 | 0x0A, 10Hz | [Mode] +0x00: MIS2DH buffer +0x80: MCU buffer ID: 0x18 or 0x19 | 0xXX, 0xXX | 0x47 | 7 | 0x00 - 0x04 | ID: 0x18 or 0x19 | X0, uint16_t value(LSB), 2bytes | Y0, uint16_t value(LSB), 2bytes | Z0, uint16_t value(LSB), 2bytes | 0xXX, 0xXX
0x19, 25Hz
0x32, 50Hz
0x64, 100Hz
0x01: Buffer overflow0
0x02: Communication error0
0x03: Device not ready0
0x04: Invalid configuration

## STOP_ACC
Request Command |  |  |  | Response Command
Opcode | Plen | P-DevID, 1b | CRC16 | Opcode | Plen | Status | P-DevID, 1b | CRC16
0x49 | 1 | ID: 0x18 or 0x19 | 0xXX, 0xXX | 0x49 | 1 | 0x00 | ID: 0x18 or 0x19 | 0xXX, 0xXX

## R_HP
Request Command |  |  |  | Response Command | [5 empty] | Parameter Name
Opcode | Plen | P-ID (ASCII, 4b) | CRC16 | Opcode | Plen | Status | P-ID (ASCII, 4b) | P-Data | CRC16
0xAA | 4 | "HP01" | 0xXX, 0xXX | 0xAA | 4 + 8 | 0x00 | "HP01" | ASCII, 8bytes | 0xXX, 0xXX | IPG_SERIAL_NUMBER
0xAA | 4 | "HP02" | 0xXX, 0xXX | 0xAA | 4 + 6 | 0x00 | "HP02" | ASCII, 6bytes | 0xXX, 0xXX | IPG_MODEL
0xAA | 4 | "HP03" | 0xXX, 0xXX | 0xAA | 4 + 24 | 0x00 | "HP03" | Binary data, 24bytes | 0xXX, 0xXX | IPG_BLE_ID
0xAA | 4 | "HP04" | 0xXX, 0xXX | 0xAA | 4 + 10 | 0x00 | "HP04" | Binary data, 10bytes | 0xXX, 0xXX | IPG_PRODUCTION_LOCATION
0xAA | 4 | "HP05" | 0xXX, 0xXX | 0xAA | 4 + 6 | 0x00 | "HP05" | ASCII, 6bytes | 0xXX, 0xXX | IPG_FW_VERSION
0xAA | 4 | "HP06" | 0xXX, 0xXX | 0xAA | 4 + 10 | 0x00 | "HP06" | ASCII, 10bytes | 0xXX, 0xXX | IPG_MANUFACTURING_DATE
0xAA | 4 | "HP07" | 0xXX, 0xXX | 0xAA | 4 + 10 | 0x00 | "HP07" | ASCII, 10bytes | 0xXX, 0xXX | IPG_IMPLANTATION_DATE
0xAA | 4 | "HP08" | 0xXX, 0xXX | 0xAA | 4 + 24 | 0x00 | "HP08" | Binary data, 24bytes | 0xXX, 0xXX | LINKED_CRC_BLE_ID
0xAA | 4 | "HP09" | 0xXX, 0xXX | 0xAA | 4 + 64 | 0x00 | "HP09" | Binary data, 64bytes | 0xXX, 0xXX | LINKED_CRC_KEY
0xAA | 4 | "HP10" | 0xXX, 0xXX | 0xAA | 4 + 6 | 0x00 | "HP10" | ASCII, 6bytes | 0xXX, 0xXX | LINKED_CRC_FW_VERSION
0xAA | 4 | "HP11" | 0xXX, 0xXX | 0xAA | 4 + 24 | 0x00 | "HP11" | Binary data, 24bytes | 0xXX, 0xXX | LINKED_PRC_BLE_ID
0xAA | 4 | "HP12" | 0xXX, 0xXX | 0xAA | 4 + 64 | 0x00 | "HP12" | Binary data, 64bytes | 0xXX, 0xXX | LINKED_PRC_KEY
0xAA | 4 | "HP13" | 0xXX, 0xXX | 0xAA | 4 + 6 | 0x00 | "HP13" | ASCII, 6bytes | 0xXX, 0xXX | LINKED_PRC_FW_VERSION
0xAA | 4 | "HP14" | 0xXX, 0xXX | 0xAA | 4 + 2 | 0x00 | "HP14" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | BLE_BROADCAST_TIMEOUT
0xAA | 4 | "HP15" | 0xXX, 0xXX | 0xAA | 4 + 2 | 0x00 | "HP15" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | BLE_IDLE_CONNECTION
0xAA | 4 | "HP16" | 0xXX, 0xXX | 0xAA | 4 + 2 | 0x00 | "HP16" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | BLE_DISCONNECT_REQUEST
0xAA | 4 | "HP17" | 0xXX, 0xXX | 0xAA | 4 + 2 | 0x00 | "HP17" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | BLE_INTERVAL
0xAA | 4 | "HP18" | 0xXX, 0xXX | 0xAA | 4 + 2 | 0x00 | "HP18" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | IMPEDANCE_TEST_INTERVAL
0xAA | 4 | "HP19" | 0xXX, 0xXX | 0xAA | 4 + 2 | 0x00 | "HP19" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | BATTERY_TEST_INTERVAL
0xAA | 4 | "HP20" | 0xXX, 0xXX | 0xAA | 4 + 2 | 0x00 | "HP20" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | MAGNET_SHUTDOWN_MIN_TIME
0xAA | 4 | "HP21" | 0xXX, 0xXX | 0xAA | 4 + 2 | 0x00 | "HP21" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | MAGNET_SHUTDOWN_MAX_TIME
0xAA | 4 | "HP22" | 0xXX, 0xXX | 0xAA | 4 + 2 | 0x00 | "HP22" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | MAGNET_WAKEUP_MIN_TIME
0xAA | 4 | "HP23" | 0xXX, 0xXX | 0xAA | 4 + 2 | 0x00 | "HP23" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | MAGNET_WAKEUP_MAX_TIME
0xAA | 4 | "HP24" | 0xXX, 0xXX | 0xAA | 4 + 2 | 0x00 | "HP24" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | MAGNET_RESET_MIN_TIME
0xAA | 4 | "HP25" | 0xXX, 0xXX | 0xAA | 4 + 2 | 0x00 | "HP25" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | MAGNET_RESET_MAX_TIME
0xAA | 4 | "HP26" | 0xXX, 0xXX | 0xAA | 4 + 2 | 0x00 | "HP26" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | BATTERY_ER_LEVEL
0xAA | 4 | "HP27" | 0xXX, 0xXX | 0xAA | 4 + 2 | 0x00 | "HP27" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | BATTERY_EOS_LEVEL
0xAA | 4 | "HP28" | 0xXX, 0xXX | 0xAA | 4 + 2 | 0x00 | "HP28" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | LANGUAGE
0xAA | 4 | "HP29" | 0xXX, 0xXX | 0xAA | 4 + 2 | 0x00 | "HP29" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | IDLE_DURATION
0xAA | 4 | "HP30" | 0xXX, 0xXX | 0xAA | 4 + 2 | 0x00 | "HP30" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | RTC_INTERRUPT_PERIOD

## W_HP
Request Command |  |  |  |  | Response Command |  |  |  | Parameter Name
Opcode | Plen | P-ID (ASCII, 4b) | P-Data | CRC16 | Opcode | Plen | Status | CRC16
0xAB | 4 + 8 | "HP01" | ASCII, 8bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | IPG_SERIAL_NUMBER
0xAB | 4 + 6 | "HP02" | ASCII, 6bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | IPG_MODEL
0xAB | 4 + 24 | "HP03" | Binary data, 24bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | IPG_BLE_ID
0xAB | 4 + 10 | "HP04" | Binary data, 10bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | IPG_PRODUCTION_LOCATION
0xAB | 4 + 6 | "HP05" | ASCII, 6bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | IPG_FW_VERSION
0xAB | 4 + 10 | "HP06" | ASCII, 10bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | IPG_MANUFACTURING_DATE
0xAB | 4 + 10 | "HP07" | ASCII, 10bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | IPG_IMPLANTATION_DATE
0xAB | 4 + 24 | "HP08" | Binary data, 24bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | LINKED_CRC_BLE_ID
0xAB | 4 + 64 | "HP09" | Binary data, 64bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | LINKED_CRC_KEY
0xAB | 4 + 6 | "HP10" | ASCII, 6bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | LINKED_CRC_FW_VERSION
0xAB | 4 + 24 | "HP11" | Binary data, 24bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | LINKED_PRC_BLE_ID
0xAB | 4 + 64 | "HP12" | Binary data, 64bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | LINKED_PRC_KEY
0xAB | 4 + 6 | "HP13" | ASCII, 6bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | LINKED_PRC_FW_VERSION
0xAB | 4 + 2 | "HP14" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | BLE_BROADCAST_TIMEOUT
0xAB | 4 + 2 | "HP15" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | BLE_IDLE_CONNECTION
0xAB | 4 + 2 | "HP16" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | BLE_DISCONNECT_REQUEST
0xAB | 4 + 2 | "HP17" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | BLE_INTERVAL
0xAB | 4 + 2 | "HP18" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | IMPEDANCE_TEST_INTERVAL
0xAB | 4 + 2 | "HP19" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | BATTERY_TEST_INTERVAL
0xAB | 4 + 2 | "HP20" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | MAGNET_SHUTDOWN_MIN_TIME
0xAB | 4 + 2 | "HP21" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | MAGNET_SHUTDOWN_MAX_TIME
0xAB | 4 + 2 | "HP22" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | MAGNET_WAKEUP_MIN_TIME
0xAB | 4 + 2 | "HP23" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | MAGNET_WAKEUP_MAX_TIME
0xAB | 4 + 2 | "HP24" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | MAGNET_RESET_MIN_TIME
0xAB | 4 + 2 | "HP25" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | MAGNET_RESET_MAX_TIME
0xAB | 4 + 2 | "HP26" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | BATTERY_ER_LEVEL
0xAB | 4 + 2 | "HP27" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | BATTERY_EOS_LEVEL
0xAB | 4 + 2 | "HP28" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | LANGUAGE
0xAB | 4 + 2 | "HP29" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | IDLE_DURATION
0xAB | 4 + 2 | "HP30" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAB | 0 | 0x00 | 0xXX, 0xXX | RTC_INTERRUPT_PERIOD

## R_SP
Request Command |  |  |  | Response Command | [5 empty] | Parameter Name V1 | Parameter Name V2
Opcode | Plen | P-ID (ASCII, 4b) | CRC16 | Opcode | Plen | Status | P-ID (ASCII, 4b) | P-Data | CRC16
0xAC | 4 | "SP01" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "SP01" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | NUMBER_OF_THERAPY_SESSIONS | NUMBER_OF_THERAPY_SESSIONS
0xAC | 4 | "SP02" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "SP02" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | PULSE_AMPLITUDE | PULSE_AMPLITUDE
0xAC | 4 | "SP03" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "SP03" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | PULSE_WIDTH | PULSE_WIDTH
0xAC | 4 | "SP04" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "SP04" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | PULSE_FREQUENCY | PULSE_FREQUENCY
0xAC | 4 | "SP05" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "SP05" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | RAMP_UP_DURATION | RAMP_UP_DURATION
0xAC | 4 | "SP06" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "SP06" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | RAMP_DOWN_DURATION | RAMP_DOWN_DURATION
0xAC | 4 | "SP07" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "SP07" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | TRAIN_ON_DURATION | TRAIN_ON_DURATION
0xAC | 4 | "SP08" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "SP08" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | TRAIN_OFF_DURATION | TRAIN_OFF_DURATION
0xAC | 4 | "SP09" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "SP09" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | CATHODE_ELECTRODE_NUMBER | SNS_CATHODE_ELECTRODE_NUMBER
0xAC | 4 | "SP10" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "SP10" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | ANODE_ELECTRODE_NUMBER | SNS_ANODE_ELECTRODE_NUMBER
0xAC | 4 | "SP11" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "SP11" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | MAX_SAFE_AMPLITUDE | VNS_CATHODE_ELECTRODE_NUMBER
0xAC | 4 | "SP12" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "SP12" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | MIN_SAFE_IMPEDANCE | VNS_ANODE_ELECTRODE_NUMBER
0xAC | 4 | "SP13" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "SP13" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX |  | MAX_SAFE_AMPLITUDE
0xAC | 4 | "SP14" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "SP14" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX |  | MIN_SAFE_IMPEDANCE
0xAC | 4 | "SP15" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "SP15" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX |  | SINE_AMPLITUDE
0xAC | 4 | "SP16" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "SP16" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX |  | MAX_SAFE_SINE_AMPLITUDE
0xAC | 4 | "SP17" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "SP17" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX |  | SINE_FREQUENCY
0xAC | 4 | "SP18" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "SP18" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX |  | VNSB_ON_DURATION
0xAC | 4 | "SP19" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "SP19" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX |  | VNSB_OFF_DURATION
0xAC | 4 | "ST11" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "ST11" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | THERAPY_SESSION_1_START | THERAPY_SESSION_1_START
0xAC | 4 | "ST12" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "ST12" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | THERAPY_SESSION_1_STOP | THERAPY_SESSION_1_STOP
0xAC | 4 | "ST21" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "ST21" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | THERAPY_SESSION_2_START | THERAPY_SESSION_2_START
0xAC | 4 | "ST22" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "ST22" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | THERAPY_SESSION_2_STOP | THERAPY_SESSION_2_STOP
0xAC | 4 | "ST31" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "ST31" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | THERAPY_SESSION_3_START | THERAPY_SESSION_3_START
0xAC | 4 | "ST32" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "ST32" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | THERAPY_SESSION_3_STOP | THERAPY_SESSION_3_STOP
0xAC | 4 | "ST41" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "ST41" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | THERAPY_SESSION_4_START | THERAPY_SESSION_4_START
0xAC | 4 | "ST42" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "ST42" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | THERAPY_SESSION_4_STOP | THERAPY_SESSION_4_STOP
0xAC | 4 | "ST51" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "ST51" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | THERAPY_SESSION_5_START | THERAPY_SESSION_5_START
0xAC | 4 | "ST52" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "ST52" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | THERAPY_SESSION_5_STOP | THERAPY_SESSION_5_STOP
0xAC | 4 | "ST61" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "ST61" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | THERAPY_SESSION_6_START | THERAPY_SESSION_6_START
0xAC | 4 | "ST62" | 0xXX, 0xXX | 0xAC | 4 + 2 | 0x00 | "ST62" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | THERAPY_SESSION_6_STOP | THERAPY_SESSION_6_STOP

## W_SP
Request Command |  |  |  |  | Response Command |  |  |  | Parameter Name V1 | Parameter Name V2
Opcode | Plen | P-ID (ASCII, 4b) | P-Data | CRC16 | Opcode | Plen | Status | CRC16
0xAD | 4 + 2 | "SP01" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | NUMBER_OF_THERAPY_SESSIONS | NUMBER_OF_THERAPY_SESSIONS
0xAD | 4 + 2 | "SP02" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | PULSE_AMPLITUDE | PULSE_AMPLITUDE
0xAD | 4 + 2 | "SP03" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | PULSE_WIDTH | PULSE_WIDTH
0xAD | 4 + 2 | "SP04" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | PULSE_FREQUENCY | PULSE_FREQUENCY
0xAD | 4 + 2 | "SP05" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | RAMP_UP_DURATION | RAMP_UP_DURATION
0xAD | 4 + 2 | "SP06" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | RAMP_DOWN_DURATION | RAMP_DOWN_DURATION
0xAD | 4 + 2 | "SP07" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | TRAIN_ON_DURATION | TRAIN_ON_DURATION
0xAD | 4 + 2 | "SP08" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | TRAIN_OFF_DURATION | TRAIN_OFF_DURATION
0xAD | 4 + 2 | "SP09" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | CATHODE_ELECTRODE_NUMBER | SNS_CATHODE_ELECTRODE_NUMBER
0xAD | 4 + 2 | "SP10" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | ANODE_ELECTRODE_NUMBER | SNS_ANODE_ELECTRODE_NUMBER
0xAD | 4 + 2 | "SP11" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | MAX_SAFE_AMPLITUDE | VNS_CATHODE_ELECTRODE_NUMBER
0xAD | 4 + 2 | "SP12" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | MIN_SAFE_IMPEDANCE | VNS_ANODE_ELECTRODE_NUMBER
0xAD | 4 + 2 | "SP13" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX |  | MAX_SAFE_AMPLITUDE
0xAD | 4 + 2 | "SP14" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX |  | MIN_SAFE_IMPEDANCE
0xAD | 4 + 2 | "SP15" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX |  | SINE_AMPLITUDE
0xAD | 4 + 2 | "SP16" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX |  | MAX_SAFE_SINE_AMPLITUDE
0xAD | 4 + 2 | "SP17" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX |  | SINE_FREQUENCY
0xAD | 4 + 2 | "SP18" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX |  | VNSB_ON_DURATION
0xAD | 4 + 2 | "SP19" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX |  | VNSB_OFF_DURATION
0xAD | 4 + 2 | "ST11" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | THERAPY_SESSION_1_START | THERAPY_SESSION_1_START
0xAD | 4 + 2 | "ST12" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | THERAPY_SESSION_1_STOP | THERAPY_SESSION_1_STOP
0xAD | 4 + 2 | "ST21" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | THERAPY_SESSION_2_START | THERAPY_SESSION_2_START
0xAD | 4 + 2 | "ST22" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | THERAPY_SESSION_2_STOP | THERAPY_SESSION_2_STOP
0xAD | 4 + 2 | "ST31" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | THERAPY_SESSION_3_START | THERAPY_SESSION_3_START
0xAD | 4 + 2 | "ST32" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | THERAPY_SESSION_3_STOP | THERAPY_SESSION_3_STOP
0xAD | 4 + 2 | "ST41" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | THERAPY_SESSION_4_START | THERAPY_SESSION_4_START
0xAD | 4 + 2 | "ST42" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | THERAPY_SESSION_4_STOP | THERAPY_SESSION_4_STOP
0xAD | 4 + 2 | "ST51" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | THERAPY_SESSION_5_START | THERAPY_SESSION_5_START
0xAD | 4 + 2 | "ST52" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | THERAPY_SESSION_5_STOP | THERAPY_SESSION_5_STOP
0xAD | 4 + 2 | "ST61" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | THERAPY_SESSION_6_START | THERAPY_SESSION_6_START
0xAD | 4 + 2 | "ST62" | Steps, uint16_t value(LSB), 2bytes | 0xXX, 0xXX | 0xAD | 0 | 0x00 | 0xXX, 0xXX | THERAPY_SESSION_6_STOP | THERAPY_SESSION_6_STOP

## R_LOG
Request Command | [9 empty] | Response Command
Opcode | Plen | Payload(The time stamp used to search the log) | [6 empty] | CRC16 | Opcode | Plen | Status | Payload | CRC16
0xAE | 7 | uint8_t Year, 1b | uint8_t Month, 1b | uint8_t Day, 1b | uint8_t Hour, 1b | uint8_t Min, 1b | uint8_t Sec, 1b | uint8_t SubSec, 1b | 0xXX, 0xXX | 0xAE | Log Length | 0x00 | Log, ASCII | 0xXX, 0xXX

## R_TD
Request Command |  |  |  | Response Command
Opcode | Plen | Payload | CRC16 | Opcode | Plen | Status | Payload(Date and Time) | [5 empty] | CRC16
0xB0 | 0 | --- | 0xXX, 0xXX | 0xB0 | 6 | 0x00 | uint8_t Year, 1b | uint8_t Month, 1b | uint8_t Day, 1b | uint8_t Hour, 1b | uint8_t Min, 1b | uint8_t Sec, 1b | 0xXX, 0xXX
READ_TIME_DATE

## W_TD
Request Command | [8 empty] | Response Command
Opcode | Plen | Payload(Date and Time) | [5 empty] | CRC16 | Opcode | Plen | Status | Payload | CRC16
0xB1 | 6 | uint8_t Year, 1b | uint8_t Month, 1b | uint8_t Day, 1b | uint8_t Hour, 1b | uint8_t Min, 1b | uint8_t Sec, 1b | 0xXX, 0xXX | 0xB1 | 0 | 0x00 | --- | 0xXX, 0xXX
WRITE_TIME_DATE

## AUTH
Request Command | [7 empty] | Response Command
Opcode | Payload Length | Payload | Payload | Payload | Payload | Payload | CRC16 | Opcode | plen | Status | Payload | CRC16
0xF0 | 32 + 32 + 32 + 6 + 4 | ECDSA signature part r, 32bytes | ECDSA signature part s, 32bytes | Hash message, 32bytes | Linked FW ver, 6bytes Unused – fill with 0x00 | Linked BLE ID, 4bytes Unused – fill with 0x00 | 0xXX, 0xXX | 0xF0 | 1 | 0x00 0x00 0x00 0x01 | User Class, 1byte 0xFF: USER_CLASS_ADMIN 0x02: USER_CLASS_CLINICIAN 0x01: USER_CLASS_PATIENT 0x00: USER_CLASS_INVALID | 0xXX, 0xXX

## R_BLE
Request Command |  |  |  | Response Command
Opcode | plen | Payload | CRC16 | Opcode | plen | Status | Payload | Payload | CRC16
0xF1 | 0 | --- | 0xXX, 0xXX | 0xF1 | 6 + 1 | 0x00 | Passkey numbers, ASCII, 6bytes | whitelist_enable, bool, 1byte | 0xXX, 0xXX
READ_BLE_ADVANCE

## W_BLE
Request Command |  |  |  |  | Response Command
Opcode | plen | Payload | Payload | CRC16 | Opcode | plen | Status | Payload | CRC16
0xF2 | 6 + 1 | Passkey numbers, ASCII, 6bytes | whitelist_enable, bool, 1byte | 0xXX, 0xXX | 0xF2 | 0 | 0x00 | --- | 0xXX, 0xXX
WRITE_BLE_ADVANCE

## AUTH_FW_IMAGE
Request Command | [5 empty] | Response Command
Opcode | plen | Payload | Payload | Payload | CRC16 | Opcode | plen | Status | Payload | CRC16
0xF3 | 32 + 32 + 32 | ECDSA signature part r, 32bytes | ECDSA signature part s, 32bytes | Hash message, 32bytes | 0xXX, 0xXX | 0xF3 | 0 | 0x00 | --- | 0xXX, 0xXX

## DOWNLOAD_FW_IMAGE
Request Command |  |  |  |  | Response Command
Opcode | plen | Payload | Payload | CRC16 | Opcode | plen | Status | Payload | CRC16
0xF4 | 4 + 128 | ImageDataOffset, uint32_t, LSB, 4bytes The offset of the data in the image | ImageData, 128bytes The data of the packet | 0xXX, 0xXX | 0xF4 | 0 | 0x00 | --- | 0xXX, 0xXX

## VERIFY_FW_IMAGE
Request Command |  |  |  | Response Command
Opcode | plen | Payload | CRC16 | Opcode | plen | Status | Payload | CRC16
0xF5 | 4 | Image Size, uint32_t, LSB, 4bytes The size of the image | 0xXX, 0xXX | 0xF5 | 0 | 0x00 | --- | 0xXX, 0xXX

## SET_START_STATE
Request Command |  |  |  | Response Command
Opcode | plen | Payload | CRC16 | Opcode | plen | Status | Payload | CRC16
0xF6 | 2 | System start state, uint16_t, LSB, 2bytes Apply after reboot 0x0100: STATE_SLEEP 0x0200: STATE_ACT 0x0201: STATE_ACT_MODE_BLE_ACT 0x0203: STATE_ACT_MODE_THERAPY_SESSION 0x0204: STATE_ACT_MODE_IMPED_TEST 0x0205: STATE_ACT_MODE_BATT_TEST 0x0208: STATE_ACT_MODE_DVT | 0xXX, 0xXX | 0xF6 | 0 | 0x00 | --- | 0xXX, 0xXX